## Slide 1: Service Presentation

## Slide 2: NanoRam 1
Handheld Raman instrument primarily use for identification/verification of materials
Originally developed from MiniRam, same probe shaft diameter and working distance
Optical bench- Similar to BTR113
Pros: Flagship handheld Raman unit, allowed access to a big market (pharma)
Cons: Software issues, sticky top panel, WinCE, laser cable is intended for data, worse performance than iRaman

## Slide 3: NanoRam 1 Details
WinCE based firmware, NOS/NID software
785nm excitation wavelength laser
Allows for library/method development, data storage/transfer/export/reports
Write firmware, back up customer data, create zip file to restore customer data, update NOS version, update library, test core modules verify system works correctly, FQC, replace top panel, core module, main PCB, FFC cables, etc.

## Slide 4: NanoRam 1 Inside
Similar to NR2 with a few differences
LED and speaker wires soldered directly to PCB
Battery slot on other end of main PCB
Wifi antenna attaches from core module to inside of top panel
Top panel ports look slightly different
Bigger probe shaft and altered optical bench

## Slide 5: NanoRam 2
Mainly intended for laboratories, pharma, base library package comes with excipients and APIs
Pros: TE Cooled, method verification, main board + spectrometer board separate, no WinCE
Cons: WinCE, too many versions of parts( main board, touch screen, firmware, etc), laser driven by FFC cable

## Slide 6: NanoRam 2 Details
Linux based firmware (v9 and v11)
NOS embedded software and NID software
785nm excitation wavelength laser
Allows for library/method development, data storage/transfer/export/reports
Write firmware, back up customer data, create zip file to restore customer data, update NOS version, update library, test core modules verify system works correctly, FQC

## Slide 7: NanoRam 2 Inside
Internal main board parts(CCW): Speaker connector, interlock, USB port with identity crisis, thinks it’s ethernet, external power port, LED connector, SD card reader, top panel 50 pin FFC cable connector, spectrometer board cable, laser board cable, lithium ion battery, TEC board cable
Bottom main board (CCW): Speaker, battery board connector, core module, barcode scanner, wifi antenna attaches to core module

## Slide 8: NR2 Hardware Compatibility
Top panel (new old), core module (4), display driver board (old revision E, old revision G, new one revision A, A and G are homologous), main PCB(1.0-03 old, 77-1.0 new wifi cable hole), OMAP 350 firmware (9-1, 9-2, 11-1, 11-2)

## Slide 9: NanoRam Mini
I’ve never seen one
Basically a TTD with a barcode scanner and NR software, allegedly
No TE cooling and possibly a software difference?
Claim of being smaller, no proof yet

## Slide 10: Tactic ID
Another class of handheld Raman system, designed for identification and verification of suspicious, dangerous, unknown, and illegal substances in the field by safety personnel
Targeted for field use for narcotics and explosives libraries
Pros: Lots of variety (GP/N/1064), advanced mixture analysis, multiple LEW options, SERS
Cons: Only matches using HQI (no p-value)

## Slide 11: Tactic ID Insides
No spectrometer TEC
Only main PCB, no spectrometer board
No barcode reader
Otherwise, fairly similar to NR2

## Slide 12: Tactic ID Hardware Compatibility
Top panels (new, old), main PCB(revision C need to chop up the chassis to upgrade from C, D added extra sensing wire to battery + added disposable lithium ion battery, E incorporated D hardware upgrades , F supported new top panel), firmware (two for old top panel, two for new top panel), core module( 1 chip vs 2 chip on core)

## Slide 13: HHEX (TacticID-1064)
Designed for field use for personnel to have non-destructive identification of unknown/dangerous/illegal substances
Pros: Relatively fast results from avoiding fluorescence, mixture analysis, customizable libraries, certified water resistant
Cons: Not good at ID’ing plant materials like marijuana, overheating problems due to deep cooling InGaAs detector, battery life, can’t be reopened,

## Slide 14: HHEX Details
Linux based firmware written on square board
TOS EX embedded and TID EX software
1064nm excitation wavelength laser (NIR)
Write firmware to core module, create factory settings and files through Filezilla, TID EX, and Prism file system, update TOS EX, update libraries and safety information, check unit functionality

## Slide 15: HHEX Insides
Almost everything is different from TTD and NR2
Different optical path, different number of pixels (512), different main PCB, different spectrometer board, different LEW and laser board, NR2 and HHEX are TEC, TTD is not
Spectrometer similar to 284

## Slide 16: Analytic Methods
HQI vs P-value: What’s the difference?
HQI: Measure of how well a query spectrum compares against a reference spectrum
  Calculated by cosine of angle between spectra
P-value: Probability of obtaining a result at least as extreme as the observed result
  Chi square test using observed + expected data
HQI suggests the best matches, while p-value tells the user the quality of each match

## Slide 17: Spectrometer Model Numbers
First letter “B” is for BW Tek
Second letter “R”/”T” denotes whether or not the unit is TE cooled
Third letter “C” indicates compact device
First number indicates type of detector
Second number indicates optical bench
Third number indicates method of communication

## Slide 18: Fibers + Probes I
FRS round-to-slit: collects light with bundle fiber and emits into slit in linear stack
  Higher throughput and SNR, same resolution
BFA: bifurcated optic fiber bundle
  BRS is version for OEM only
TFA: trifurcated fiber optic bundle (6v1 config)
FRP: fiber reflectance probe, measures diffuse/specular reflectance
  Normal to surface, measures specular reflectance
  At an angle, measures diffuse reflectance

## Slide 19: Fibers + Probes II
FDP: fiber optic dip probe
  Similar to reflection probe, but cavity fills with liquid sample and undergoes transflectance
FRC100: fiber optic near cosine corrector
  Diffuser that is more efficient and sensitive than integrating spheres over a broad spectral range
BAT100: manual attenuator, used by alignment technicians at BWTEK

## Slide 20: Fibers+ Probes III
SRR: Spectral reflectance reference
  Pure white material used for calibration for reflectance measurements
CLA-SMA: fiber optic collimating lens
CS-SMA/C-CS/CS-SMAR: some kind of lens mount?

## Slide 21: Raman Probe
Used to measure the inelastic scattering of light off of a sample
Band-pass used to make sure the signal from the excitation source is as pure as possible
Long-pass filter used to remove Rayleigh scattering and laser excitation
Dichroic filter used to redirect backscattered light to collection fiber

## Slide 22: Raman Portable Systems
STRam used for identifying materials through barriers or packaging
QTRam used for analyzing blend and content uniformity of products (gels, tablets, capsules)
iRaman has Sony ILX detector
iRaman Plus is iRaman with back-thinned Hamamatsu detector
iRaman Pro is a Plus with deeper cooling
iRaman Prime is a Pro with high throughput linear stacked fibers and bigger slit
STRam is Prime with special tip with ST tech
QTRam is STRam with transmission probe

## Slide 23: Lasers
BWF1: Full laser system with wide range of wavelengths offered (600-1600nm), class III laser with high brightness and fiber coupling, laser power up to 450mW
BWF2: Full laser system with even wider wavelength range offered (445-2000nm), higher powered laser up to 30W, fiber delivered (not coupled)
CleanLaze has a very narrow linewidth from using a grating to filter out side lobes and maintain constant central wavelength

## Slide 24: CCD Spectrometers
Exemplar (???)
Exemplar LS (???)
Exemplar Plus (BTC655)
Exemplar Plus LS (BTC645N)
Exemplar Pro (Exemplar Pro)
Glacier Pro (BTC675)
Glacier X (BTC112E)- intended for standalone, others usually go in some version of iRaman?
Glacier T (BTC162E)- goes in iRaman original

## Slide 25: InGaAs Spectrometers
Sol 1.7 (?)
Sol 2.2A (?)
Sol 2.6 (BTC 263)
Big range out to mid IR, higher noise detectors
Mid IR good for organic molecules

## Slide 26: Annual Certification
Proof for the customer that the unit is calibrated correctly and verifies authenticity of future data
Keeps the FDA from locking up Big Pharma